## Slide 1: Quality Laboratory Results
Presenting: Kristen Durie
Chemistry Quality Assurance Officer
New York State Food Laboratory

## Slide 2: Located in Albany, New York

## Slide 3: Food Laboratory Background:
Year | # Samples | Total # Recalls | Class 1 | Class II | Class III
2012 | 23,268 | 308 | 34 Imported | 176 Imported | 19 Imported
 |  |  | 10 Domestic | 61 Domestic | 9 Domestic
2013 | 16,312 | 292 | 31 Imported | 161 Imported | 24 Imported
 |  |  | 12 Domestic | 51 Domestic | 10 Domestic
2014 (to date) | 7,903 | 85 | 12 Imported | 40 Imported | 12 Imported
 |  |  | 3 Domestic | 14 Domestic | 4 Domestic

## Slide 4: How is our Quality Ensured?

## Slide 5: Management System:
ISO 17025 Accredited
  Document Control
    Methods
      All Methods are revision controlled and distributed
    Work Instructions
      Revision controlled and distributed
      Written in addition to and to enhance methods
    Forms
      Revision controlled and made available at all times to analysts

## Slide 6: Management System:
Control of Records
  Records include but are not limited to:
    Bench Worksheets
    Instrument Data
    Worksheets
    Personal Laboratory Notebooks
    Databases
    Sample information

## Slide 7: Management System:
Internal Audits
  Each member of the Quality Assurance Unit (QAU) receives training to become an auditor
  Each General Lab Procedure, SOP (Instrument and Method), Standard Instruction and Corrective action are audited on a 2 year cycle.
    By auditing each of the above the QAU have an opportunity to ensure that the method is current and being followed correctly.
External Audits
  Every 2 years A2LA audits the laboratory

## Slide 8: Management System:
Corrective Actions
  Audits
  Proficiency test failures
  Control chart data

## Slide 9: Management System:
Equipment
  All fridges, freezers, autoclaves, incubators, and ovens are around the clock monitored using a temperature monitoring program Tempsys by Checkpoint
  All instruments are maintained in house with regular maintenance as well as a scheduled preventative maintenance performed by the manufacturer
  All pipettes are calibrated every 6 months using an Artel System
  Every critical balance is internally calibrated every day and documented by Lab X software

## Slide 10: Microbiology:
Every run has a media blank, a positive and negative control prepped and analyzed with the samples
Participates in ~7 proficiency tests a year (Listeria, L.mono, Salmonella, Campylobacter, Staph, and E.coli)
Proficiency tests are broken up between analysts. Each analyst will perform a different part of the analysis per PT.
All food testing samples are done in duplicate

## Slide 11: Microbiology:
Training
  Micro analysts are sent to outside training when relevant. Most recently a bench microbiologist attended PFGE training and the food microbiology supervisor attended a FERN training covering Anthrax and Plague
  In-house training
    Training Check lists are used
    Control run is performed
    Passing proficiency test results

## Slide 12: Microbiology: Method Validation
Validate 3 Campylobacter species detection on the BAX (PCR Detector)
  Four 25 mL milk samples were set up using the same milk sample.
  One was blank matrix, and the other 3 were inoculated with a different Campylobacter control at a reasonable challenge level.
  Samples will be enriched and incubated as per MICRO-MTH-321, including standard cultural controls.
  Samples will be extracted and run on the BAX, and plated to demonstrate recoverability.
Sample Names/ Species of Campy:
M-Uninoc – uninoculated
M-CJ – C. jejuni spike
M-CC – C. coli spike
M-CL – C. lari spike
Sample Levels:
CJ – 1.7 cfu/mL
CC – 0.3 cfu/mL
CL – 1.2 cfu/mL

## Slide 13: Pesticide Data Program (PDP):
Completion of 1 proficiency test per year
With every run there are: (USDA Mandated)
  Process Controls (Surrogate)
  Matrix Blanks
  Reagent Blanks
  Fortification Samples
Above and beyond the USDA Mandates:
  Spike with all compounds (~200) with each sample set (<25 samples)
  Full set of Calibration Standards

## Slide 14: Pesticide Data Program (PDP):
  PDP analysts are sent to outside training when relevant. Most recently a number of PDP analysts were sent to PittCon to attend chromatography and mass spectroscopy courses.
  In-house training
    Training Check lists are used
    Training is done concurrently with sample preparation
    Passing proficiency test results

## Slide 15: Chemistry:
Similar to PDP
  Reagent Blanks
  Matrix Blanks
  Fortification Samples
  Standard Curve is generated for all quantitative analyses
  Standards purchased from a Guide 34 Accredited Establishment (when possible)
  Participate in ~40 proficiency tests a year
Chemistry Specific Quality Measures
  Reference Materials (when possible)
  Quality Control (QC) Samples (when possible)
  When a QC or reference material not available a duplicate analysis is done

## Slide 16: Chemistry Training:
  Chemists are sent to outside training when relevant. Most recently a number of Chemists were sent to PittCon to attend chromatography and mass spectroscopy courses.
  Training is periodically brought in house for Chemists as well. The most recent training was provided by the Waters Corporation.
  In-house training
    Training Check lists are used
    Training is done concurrently with sample preparation
    Passing proficiency test results

## Slide 17: Chemistry and PDP Method Validation:
Amantadine in Chicken Jerky Dog Treats
Infant Formula Commodity Validation
Validation: In Clean Matrix 2 Limit of Detection (LOD) Spiked Samples 3 Limit of Detection (LOQ) Spiked Samples 3 x5 LOQ Spiked Samples 3 x10 LOQ Spiked Samples

## Slide 18: In Closing:
Management System
Training
System Checks

## Slide 19: Thank you